OX863 Metrix

________________________________________________________________________________

Introduction:

This instrument driver provides programming support for OX863 Metrix.

It contains functions for opening, configuring, taking measurements from, and

closing the instrument.

________________________________________________________________________________

Assumptions:

To successfully use this module, the following conditions must be met:

For GPIB instrument drivers:

-  the instrument is connected to the GPIB.

-  the GPIB address supplied to the initialize function must match the

GPIB address of the instrument.

For VXI instrument drivers:

-  the instrument is installed in the VXI mainframe and you are

using one of the following controller options:

Embedded controller

MXI

MXI2

GPIB-VXI

-  the logical address supplied to the initialize function must

match the logical address of the instrument.

For RS-232 instrument drivers:

-  the instrument is connected to the RS-232 interface.

-  the COM port, baud rate, parity, and timeout supplied to the

initialize function must match the settings of the instrument.

________________________________________________________________________________

Error and Status Information:

Error codes and warnings are returned as the return value of each instrument

driver function. A program should examine this value after each call to an

instrument driver function to determine if any error occurred.  The general

meaning of the return value is as follows:

Value                   Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

Possible error codes and their meanings are listed with the corresponding

instrument driver function.

________________________________________________________________________________

How To Use This Document:

This document is intended to be used as a programming reference manual.

It describes each function in the

OX863 Metrix

instrument. The functions are listed in alphabetical order with a

description of the function, C syntax of the function, a description of each

parameter, and possible error codes.

________________________________________________________________________________

Function Tree Layout:

OX863 Metrix                           Function Name:

Initialize                                  ox863_init

Instrument Setup                            ox863_setup

ReadOut Setup                               ox863_readout

Configure

Autoset                                  ox863_autoset

Display settings                         ox863_display

Channel settings                         ox863_channel

Time base                                ox863_timebase

Trigger                                  ox863_trigger

Read-Out

Measure                                  ox863_measure

Standard TV                              ox863_standard

Line TV                                  ox863_line

Close                                       ox863_close

________________________________________________________________________________

OX863 Metrix

Instrument name: OX863 Metrix

Description: This instrument module provides

RS232 support for OX863.

The module contains the followings

functions:

(1) Initialize

Initialize the instrument module

with default values and lock the

front panel.

(2) Instrument setup

(3) Instrument Configuration

Class of functions for configuring

the following parameters:

* Autoset

* Display settings

* Channel settings

* Time Base

* Trigger

(4) Instrument Read-Out options

Class of functions for configuring

the following parameters:

* Cursors measures

* Standard TV

* Line TV

(5) Close

Close the instrument module and

unlock the front panel.

________________________________________________________________________________

The following functions are in alphabetical order.

________________________________________________________________________________

ox863_autoset

int ox863_autoset (void);

Purpose

Pannel name : Autoset

Description : This function performs an autoset of the

instrument.

Return Value

Control label  :  Error

Description    :  Displays the result of the

function call.

Variable type  :  integer

Variable range :

0  :  Success

230  :  Error writing to instrument

232  :  Instrument not initialized

Others  :  See rs232err values

________________________________________________________________________________

ox863_channel

int ox863_channel (int channel, int coupling, int range, int invertCH2,

int probe);

Purpose

Panel name  :   Channel settings

Description :   Sets the Channel , its

coupling and volts/division .

Sets the Invert CH2 .

Sets the Probe .

Coupling :  AC DC GND

Volts/Div:  1  mV/Div  to 20 V/Div

with 1-2-5 sequences .

Invert CH2 : Off On

Probe : x1, x10, x100

Parameter List

channel

Variable Type       int

Control label  : Channel number

Description    : Specifies the channel of the

oscilloscope to configure.

Variable type  : integer

Variable range :    0: CH1 (default value)

1: CH2

coupling

Variable Type       int

Control label  : Coupling

Description    : Set the coupling for the

selected channel.

Variable type  : integer

Variable range :     0 :  AC (default value)

1 :  DC

2 :  GND

range

Variable Type       int

Control label  : Range

Description    : Set the range (V/Div)

of the specified channel

Variable type  : integer

Variable range :    10   :   2  mV/Div

9   :   5  mV/Div

8   :   10 mV/Div

7   :   20 mV/Div

6   :   50 mV/Div

5   :   0.1 V/Div  (default value)

4   :   0.2 V/Div

3   :   0.5 V/Div

2   :   1   V/Div

1   :   2   V/Div

0   :   5   V/Div

invertCH2

Variable Type       int

Control label  : Invert CH2

Description    : Inverts Channel 2

Variable type  : integer

Variable range :    0 : Off  (default value)

1 : On

probe

Variable Type       int

Control label  : Probe

Description    : Selects the probe of the specified channel

Variable type  : integer

Variable range : 0 - x 1  (Default Value)

1 - x 10

2 - x 100

Return Value

Control label  :  Error

Description    :  Displays the result of the

function call.

Variable type  :  global integer

Variable range :

0  :  Success

-1  :  Channel value out of range

-2  :  Coupling value out of range

-3  :  Range value out of range

-4  :  Invert CH2 value out of range

-5  :  Coeff Probe value out of range

230   :  Error writing to instrument

232   :  Instrument not initialized

Others :  See rs232err values

________________________________________________________________________________

ox863_close

int ox863_close (void);

Purpose

Pannel name : Close

Description   : Close the instrument module

and unlock the front panel.

NOTE: The instrument must be re-

initialized to use it again.

Return Value

Control label  : Error

Description    : Displays the result of the

function call.

Variable type  : integer

Variable range :

0 :    Success

221 :    Unable to close instrument

232 :    Instrument not initialized

Others  :    See rs232err values

________________________________________________________________________________

ox863_display

int ox863_display (int verticalMode, int BWL);

Purpose

Panel name  :   Display settings

Description :   This panel is usefull to

configure the Vertical mode

and to control the Beam Find .

Vertical Mode : CH1 CH2 ALT CHOP ADD XY

Bwl           : Off On

Parameter List

verticalMode

Variable Type       int

Control label  : Vertical Mode

Description    : Sets the vertical mode

Variable type  : integer

Variable range : 0 : CH1 (default value)

1 : CH2

2 : ALT

3 : CHOP

4 : ADD

5 : XY

BWL

Variable Type       int

Control label  : BWL

Description    : This control configures the

BWL mode : On or Off .

Variable type  : integer

Variable range :    0 : Off (default value)

1 : On

Return Value

Control label  :  Error

Description    :  Displays the result of the

function call.

Variable type  :  integer

Variable range :

0  :       Success

-1  :       Vertical Mode out of range

-2  :       Bwl out of range

230   :  Error writing to instrument

232   :  Instrument not initialized

Others :  See rs232err values

________________________________________________________________________________

ox863_init

int ox863_init (int RS232portnumber);

Purpose

Panel name : Initialize

Description : Opens the instrument

with the following RS232

configuration :

9600 Bauds,8 bits,no parity

1 stop bit, RTS/CTS

and set the instrument to

default values.

Parameter List

RS232portnumber

Variable Type       int

Control label  : RS232 port number

Description    : Defines the RS232 port number

Variable type  : integer

Variable range : 1 to 8

Default value : 1

Return Value

Control label  :  Error

Description    :  Displays the result of the

function call.

Variable type  :  global integer

Variable range :

0   :  Success

-1   :  Invalid port number RS232

-2   :  Invalid Lock front panel number.

223   :  Instrument identification

query failed.

230   :  Error writing to instrument

Others   :  See rs232err

________________________________________________________________________________

ox863_line

int ox863_line (int lineTV);

Purpose

Panel name  :   Line TV

Description :   This panel sets the line TV value

Parameter List

lineTV

Variable Type       int

Control label  : Line TV

Description    : Sets line TV value

Variable type  : integer

Variable range : 1 to Standard TV

Return Value

Control label  :  Error

Description    :  Displays the result of the

function call.

Variable type  : integer

Variable range :

0  :     Success

-1  :     Line TV value out of range

230  :     Error writing to instrument

232  :     Instrument not initialized

Others  :     See rs232err values

________________________________________________________________________________

ox863_measure

int ox863_measure (int measure, int cursor1, int cursor2, int cursor3);

Purpose

Panel name  :   Measure

Description :   This panel sets the cursors on the

display :

- No measure

- Volt measure

- Time measure

- Phase measure

Parameter List

measure

Variable Type       int

Control label  : Measure

Description    : Set the cursors

Variable type  : integer

Variable range :    0 : None  (default value)

1 : Vertical

2 : Horizontal

3 : Phase

cursor1

Variable Type       int

Control label  : Cursor 1

Description    : Sets cursor 1

Variable type  : integer

Variable range : 0 to 150 in vertical measure

0 to 230 in horizontal and phase measure

cursor2

Variable Type       int

Control label  : Cursor 2

Description    : Sets cursor 2

Variable type  : integer

Variable range : 0 to 150 in vertical measure

0 to 230 in horizontal and phase measure

cursor3

Variable Type       int

Control label  : Cursor 3

Description    : Sets cursor 3

Variable type  : integer

Variable range : 0 to 230 in horizontal and phase measure

Return Value

Control label  :  Error

Description    :  Displays the result of the

function call.

Variable type  : integer

Variable range :

0  :     Success

-1  :     Measure value out of range

-2  :     Cursor 1 value out of range

-3  :     Cursor 2 value out of range

-4  :     Cursor 3 value out of range

230  :     Error writing to instrument

232  :     Instrument not initialized

Others  :     See rs232err values

________________________________________________________________________________

ox863_readout

int ox863_readout (char filename[]);

Purpose

Panel name  :   Instrument ReadOut Setup

Description :   This panel is usefull to

Save the instrument ReadOut

Setup to a file .

This file is in ASCII format and

contains the values of the

functions listed below :

Function                  Parameter         EOT

Remote Mode       : 0x9f   0x30 to 0x31      0x04

Read Out Mode     : 0xa0   0x30 to 0x31      0x04

Probe CH1         : 0xa1   0x30 to 0x32      0x04

Probe CH2         : 0xa2   0x30 to 0x32      0x04

Measure           : 0xa3   0x30 to 0x33      0x04

Cursor 1          : 0xa5   0x30 to 0xc6      0x04

or  : 0xa5   0x1f to 0xff      0x04

Cursor 2          : 0xa6   0x30 to 0xc6      0x04

or  : 0xa6   0x1f to 0xff      0x04

Cursor 3          : 0xa7   0x30 to 0xc6      0x04

or  : 0xa7   0x1f to 0xff      0x04

Standard TV       : 0xa8   0x30 to 0x93      0x04

and : 0xa9   0x35 to 0x3c      0x04

Line TV           : 0xaa   0x30 to 0x93      0x04

and : 0xab   0x35 to 0x3c      0x04

Parameter List

filename

Variable Type       char []

Control label  : File name

Description    : Defines the File name which

contains the Instrument

Setup.

Variable type  : string

Variable range : any valid string

(use \\\ with Quick C in case

of complete pathname :

Example  : "c:\\setup1.cfg")

Return Value

Control label  :  Error

Description    :  Displays the result of the

function call.

Variable type  :  global integer

Variable range :

0  :  Success

-1  :  Setup value out of range

-2  :  File doesn't exist

230   :  Error writing to instrument

231   :  Error reading from instrument

232   :  Instrument not initialized

Others :  See rs232err values

________________________________________________________________________________

ox863_setup

int ox863_setup (int setup, char filename[]);

Purpose

Panel name  :   Instrument Setup

Description :   This panel is usefull to

Save or Recall the instrument

Setup to or from a file .

This file is in ASCII format and

contains the values of the

functions listed below :

Function                  Parameter         EOT

Vertical Mode     : 0x64   0x30 to 0x35      0x04

Sensibility CH1   : 0x60   0x30 to 0x3D      0x04

Sensibility CH2   : 0x61   0x30 to 0x3D      0x04

Coupling    CH1   : 0x6a   0x30 to 0x32      0x04

Coupling    CH2   : 0x6c   0x30 to 0x32      0x04

Invert CH2        : 0x6e   0x30 to 0x31      0x04

Time Base Range A : 0x62   0x30 to 0x43      0x04

Time Base x10     : 0x71   0x30 to 0x31      0x04

Source   Trigger  : 0x68   0x30 to 0x34      0x04

Auto     Trigger  : 0x73   0x30 to 0x31      0x04

Coupling Trigger  : 0x69   0x30 to 0x35      0x04

Peak     Trigger  : 0x70   0x30 to 0x31      0x04

Slope    Trigger  : 0x72   0x30 to 0x31      0x04

Delayed  Trigger  : 0x65   0x30 to 0x32      0x04

Write Config      : 0x77   0x30              0x04

Time Base Range B : 0x63   0x30 to 0x43      0x04

Bwl               : 0x6f   0x30 to 0x31      0x04

B trig            : 0x66   0x30 to 0x31      0x04

Trace sep         : 0x67   0x30 to 0x3f      0x04

Parameter List

setup

Variable Type       int

Control label  :  Configuration

Description    :  Selects the Save or Recall

Instrument Setup function .

Variable type  : integer

Variable range :     0 : Save (default value)

1 : Recall

filename

Variable Type       char []

Control label  : File name

Description    : Defines the File name which

contains the Instrument

Setup.

Variable type  : string

Variable range : any valid string

(use \\\ with Quick C in case

of complete pathname :

Example  : "c:\\setup1.cfg")

Return Value

Control label  :  Error

Description    :  Displays the result of the

function call.

Variable type  :  global integer

Variable range :

0  :  Success

-1  :  Setup value out of range

-2  :  File doesn't exist

230   :  Error writing to instrument

231   :  Error reading from instrument

232   :  Instrument not initialized

Others :  See rs232err values

________________________________________________________________________________

ox863_standard

int ox863_standard (int standardTV);

Purpose

Panel name  :   Standard TV

Description :   This panel sets the standard TV value

Parameter List

standardTV

Variable Type       int

Control label  : Standard TV

Description    : Sets standard TV value

Variable type  : integer

Variable range : 525 to 1250

Return Value

Control label  :  Error

Description    :  Displays the result of the

function call.

Variable type  : integer

Variable range :

0  :     Success

-1  :     Standard TV value out of range

230  :     Error writing to instrument

232  :     Instrument not initialized

Others  :     See rs232err values

________________________________________________________________________________

ox863_timebase

int ox863_timebase (int x10, int rangeA, int delay, int btrig, int rangeB,

int tracesep);

Purpose

Panel name  :   Time base

Description :   This panel is usefull to set

the Time base parameters of

the oscilloscope :

x10       :  Off On

Range A   :  100 ms/Div to 50nS/Div

with 1-2-5 sequences

Delayed TB:  Off Search Delayed

B trig    :  Off On

Range B   :  100 ms/Div to 50nS/Div

with 1-2-5 sequences

Trace sep :  O step to 15 steps

Parameter List

x10

Variable Type       int

Control label  : x10

Description    : This control configures the

horizontal x10 expanded mode

On or Off .

Variable type  : integer

Variable range :    0 : Off  (default value)

1 : On

rangeA

Variable Type       int

Control label  : Range

Description    : Set the timebase Range A of

the oscilloscope .

Variable type  : integer

Variable range :    0   :   100 mS

1   :    50 mS

2   :    20 mS

3   :    10 mS

4   :     5 mS

5   :     2 mS

6   :     1 mS

7   :   500 S

8   :   200 S

9   :   100 S

10  :    50 S

11  :    20 S (default value)

12  :    10 S

13  :     5 S

14  :     2 S

15  :     1 S

16  :   0.5 S

17  :   0.2 S

18  :   0.1 S

19  :    50 nS

delay

Variable Type       int

Control label  : Delay

Description    : This control configures the

delayed time base state.

Variable type  : integer

Variable range :    0 : Off  (default value)

1 : Alterned

2 : Delay

btrig

Variable Type       int

Control label  : B trig

Description    : Set the B trig mode

Off or On .

Variable type  : integer

Variable range :   0 : Off (default value)

1 : On

rangeB

Variable Type       int

Control label  : Range

Description    : Set the timebase Range B of

the oscilloscope .

Variable type  : integer

Variable range :    0   :   100 mS

1   :    50 mS

2   :    20 mS

3   :    10 mS

4   :     5 mS

5   :     2 mS

6   :     1 mS

7   :   500 S

8   :   200 S

9   :   100 S

10  :    50 S

11  :    20 S  (default value)

12  :    10 S

13  :     5 S

14  :     2 S

15  :     1 S

16  :   0.5 S

17  :   0.2 S

18  :   0.1 S

19  :    50 nS

tracesep

Variable Type       int

Control label  : Trace sep

Description    : Set the number of step for

the Trace Sep mode.

Variable type  : integer

Variable range : 0 : 0 step  (default value)

1 : 1 step

2 : 2 steps

3 : 3 steps

4 : 4 steps

5 : 5 steps

6 : 6 steps

7 : 7 steps

8 : 8 steps

9 : 9 steps

10 : 10 steps

11 : 11 steps

12 : 12 steps

13 : 13 steps

14 : 14 steps

15 : 15 steps

Return Value

Control label  :  Error

Description    :  Displays the result of the

function call.

Variable type  : integer

Variable range :

0   :    Success

-1   :    x10 value out of range

-2   :    Range A value out of range

-3   :    Delay value out of range

-4   :    B trig value out of range

-5   :    Range B value out of range

-6   :    Trace sep value out of range

-7   :    Vertical Mode = XY

-8   :    Range B > Range A

230  :  Error writing to instrument

232  :  Instrument not initialized

Others  :  See rs232err values

________________________________________________________________________________

ox863_trigger

int ox863_trigger (int ptp, int slope, int auto, int source, int coupling);

Purpose

Panel name  :   Trigger

Description :   This panel is usefull to

configure the following

trigger parameters:

- Peak to Peak : on/off

- Slope : + / -

- Auto  : on/off

- Source : CH1,CH2,LINE,

ALT,EXT

- Coupling : AC,DC,LFR,

HFR,ALL,CPT

Parameter List

ptp

Variable Type       int

Control label  : Ptp

Description    : Set the trigger Peak to Peak

mode Off or On .

Variable type  : integer

Variable range :    0 : Off

1 : On  (default value)

slope

Variable Type       int

Control label  : Slope

Description    : Set the trigger slope

Positive or negative

Variable type  : integer

Variable range :    0 : Positive  (default value)

1 : Negative

auto

Variable Type       int

Control label  : Auto

Description    : Sets the trigger Auto mode

Off or On .

Variable type  : integer

Variable range :    0 : Off

1 : On  (default value)

source

Variable Type       int

Control label  : Source

Description    : Sets the trigger Source .

Variable type  : integer

Variable range : 0 : CH1 (default value)

1 : CH2

2 : ALT

3 : LINE

4 : EXT

When Trigger Source LINE is configured, Trigger

Coupling is disabled.

coupling

Variable Type       int

Control label  : Coupling

Description    : Sets the trigger Coupling .

Variable type  : integer

Variable range : 0 : DC  (default value)

1 : AC

2 : LFR

3 : HFR

4 : ALL

5 : CPT

When Trigger Source LINE is configured, Trigger

Coupling is disabled.

Return Value

Control label  :  Error

Description    :  Displays the result of the

function call.

Variable type  : integer

Variable range :

0  :     Success

-1  :     Ptp value out of range

-2  :     Slope value out of range

-3  :     Auto value out of range

-4  :     Source value out of range

-5  :     Coupling value out of range

-6  :     Vertical Mode = XY

230  :     Error writing to instrument

232  :     Instrument not initialized

Others  :     See rs232err values